# Introduction and P1 research purpose

RQ1 asks which B- and X-point localisation approaches are supported by available references and what contributes to their variability and failure. The answer is conditional: no single definition is reliable across every reference type, dataset and physiological condition.

HeartCycle B8 had the lowest observed B-point median absolute ECHO AVO error, while X12 combined a relatively low ECHO AVC median error with a smaller tail than several alternatives. These are reference-specific observations, not evidence of universal superiority. ReBeat expert annotations produced different patterns and are not equivalent to independent ECHO valve events. Coverage changes the interpretation: X4 failed on 48.89% of HeartCycle attempts and on about 31–47% of IITB and ReBeat attempts. IITB B1/B7 failure increased after exercise even though complete-case timing disagreement decreased.

The all-method morphology analysis linked waveform sharpness, slope, amplitude, signal quality, competing features and beat-to-beat variability to reference error, failure or disagreement in specific method-dataset combinations. The effects were heterogeneous and exploratory. They support operational reliability boundaries, not physiological causation.

## What P1 establishes

Reference correspondence, failure burden, disagreement and pairwise agreement are separate outcomes and must not be collapsed into one score.

HeartCycle B8 had the lowest observed participant median absolute B error, 12.59 ms, with signed median +11.47 ms and P95 34.21 ms; no paired superiority test was claimed.

HeartCycle X12 had the lowest observed participant median absolute X error, 36.55 ms; X2 had a similar median but a 409.40 ms P95 tail.

Common-valid X analyses retained only 371 of 691 HeartCycle reference-eligible beats, so lower disagreement among retained beats cannot be interpreted as improved reliability.

Frozen failure paths are reproducible. For X4, a third-derivative minimum at C is rejected by the strict C < X < next-R rule.

The contribution beyond earlier B-only rankings is the combined B/X reference audit, denominator-complete failure analysis, matched-condition analysis and six-feature morphology evidence.

Scope: 10 frozen B definitions and 13 frozen X definitions across HeartCycle, IITB-ICG and ReBeatICG. Canonical algorithms and calculations were not changed during this consolidation.

# Datasets references algorithms and methods

The analysis used strict common samples for reference comparisons, participant-level summaries for inference and separate failure denominators. A failed detection was never assigned a numerical timing error. ECHO AVO and AVC are valve-event measurements; ReBeat labels are expert ICG annotations; IITB has no independent B/X timing reference.

| Dataset | Conditions | Participants | Beats | Reference | Permitted claim |
| --- | --- | --- | --- | --- | --- |
| HeartCycle | Overall | 17 | 1,706 | ECHO AVO and AVC | Reference correspondence, failure and disagreement |
| IITB-ICG | Rest and post-exercise | 18 matched | 2,022 / 1,737 | None | Failure and disagreement only |
| ReBeatICG | Baseline and cognitive workload | 24 matched | 8,403 / 8,492 | Expert ICG; 1,842 matched rows | Expert-label correspondence, failure and disagreement |

## Frozen definitions and statistics

The B catalogue covers local minima, second- and third-derivative rules, amplitude thresholds and an R-to-C proportional rule. The X catalogue covers post-C minima, derivative extrema and crossings, downstroke thresholds, an RR-adaptive minimum and a metadata-gated Weissler formula. Exact parameters and implementation versions are in the canonical registry. All 23 definitions remain represented in the 115-row supplementary results table.

Accuracy: participant median absolute detected-minus-reference timing error, signed error and P90/P95 absolute error on identical common-reference beats.

Failure: missed plus invalid detections divided by attempted detections; upstream ineligibility is reported separately.

Disagreement: within-beat all-method timing range and definition-specific distance from the peer median on common-valid beats.

Condition change: stress minus baseline/rest for matched participants, with participant-bootstrap 95% intervals.

Morphology: six predeclared features, participant-aware effects and within-family Benjamini-Hochberg labels; strata were defined without access to outcomes.

Agreement: participant-balanced Bland-Altman bias and limits for related method pairs B4/B5 and X5/X6.

## Audit outcome

The 115 result rows, 514,280 method-beat records, reference formulas, common sets and denominators reconcile. Nineteen regression tests pass. No defect was found in the canonical calculations. The scientific completion adds a matched ReBeat expert-reference condition comparison, feature-specific synthesis of the full morphology screen, explicit X4 logic and an unavailable physiologically-implausible field. No clinical timing tolerance was invented.

# Reference accuracy and error distributions

Question: which definitions correspond most closely to the available reference events, and is performance consistent across participants? HeartCycle comparisons use 17 participants and strict common-reference samples of 986 B beats and 371 X beats. ReBeat expert comparisons use uniquely matched annotation rows and are interpreted separately.

Figure 2A. HeartCycle B-point reference accuracy. Each grey point is one participant's median absolute detected-minus-ECHO AVO error; black diamonds and bars show the across-participant median and IQR. B8 had the lowest observed participant median absolute error, 12.59 ms. B9 showed less signed bias (-5.82 ms) but a wider P95 tail of 57.73 ms. No paired inferential comparison established universal superiority.

X result. X12 had the lowest participant median absolute ECHO AVC error (36.55 ms; signed median +31.68 ms; P95 78.94 ms). X2 had a 41.45 ms median but a 409.40 ms P95. The full X panel and ReBeat B/X panels remain in the supplementary collection. ReBeat baseline participant medians were 4.00 ms for B1, 5.80 ms for B7 and 4.00 ms for X12/X2, but these are expert-annotation comparisons.

## Signed error distributions and tails

Median absolute error alone hides whether a definition is systematically early or late and whether a small number of beats produces extreme error. Signed error is detected time minus reference time. Negative values are early and positive values are late. All X algorithms below use the same 371 strict common-reference HeartCycle beats and common scales.

Figure 3. HeartCycle X-point signed errors against ECHO AVC. The common binning separates early and late detection while retaining complete tails. X3 and X9 were systematically early (signed medians -144.72 and -145.73 ms). X2 and X8 contain late tail clusters despite moderate central errors. X12 combined the lowest observed participant median absolute error with a smaller P95 than X2. No tolerance-exceedance percentage is shown because no clinical or analytical threshold was approved.

B distributions were narrower but still definition-dependent. B8 combined a +11.47 ms signed median with a 34.21 ms P95; B2 had a 44.01 ms participant median absolute error and 65.41 ms P95. This changes the decision from choosing a method by mean error to evaluating central correspondence, direction, tails and failure together.

# Detection success failure and variability

Question: how often does each frozen definition fail, and what recorded rule produced the failure? Failure is missed plus invalid divided by attempted detections. Upstream-ineligible records are separate and do not inflate the attempted-failure denominator.

Figure 4. HeartCycle X-point failure. Each method's participant-median failure and bootstrap interval are shown with pooled status composition. X4 was invalid on 834 of 1,706 attempts (48.89%); X8 missed 24 of 1,706 attempts (1.41%). The result establishes operational dependability, not timing accuracy.

| Method/context | Attempted failure | Verified primary reason | Interpretation |
| --- | --- | --- | --- |
| HeartCycle B1 | 195/1,705 (11.44%) | No qualifying local minimum | Failure burden accompanies otherwise moderate ECHO error |
| HeartCycle B7 | 230/1,705 (13.49%) | No crossing: 212; boundary/order: 18 | Highest HeartCycle B failure |
| HeartCycle X4 | 834/1,706 (48.89%) | X sample not strictly between C and next R | Minimum can be at C and is then rejected |

Across all datasets, 15,788 of 514,280 method-beat rows were failed or upstream-ineligible. The mutually exclusive primary reasons were 466 upstream dependency/ineligibility, 182 unusable windows, 4,520 no qualifying feature or crossing and 10,620 rejected boundary/order constraints. Unresolved count: zero. These are verified software execution paths, not physiological causes.

## Algorithm disagreement and complete case coverage

Question: how far apart are definitions on the same beat and participant? The outcome is timing disagreement, not ground-truth error. Common-valid analysis includes only beats where every method in the point-type comparison produced a valid result.

Figure 5. HeartCycle participant X-point disagreement. Each coloured dot is one participant's median within-beat timing range across all 13 X definitions. The overall participant median was about 260 ms, with strong heterogeneity. Only 852 of 1,706 beats were common-valid for this all-method display, so the plotted distribution is a success-only subset.

| Dataset/condition | B common-valid | X common-valid | Interpretation |
| --- | --- | --- | --- |
| HeartCycle overall | 1,439/1,705 | 852/1,706 | X coverage is 49.9% |
| IITB rest | 1,152/1,998 | 1,329/2,022 | Different eligible denominators |
| IITB post-exercise | 1,002/1,723 | 981/1,737 | Failure increased while retained-beat range fell |
| ReBeat baseline | 8,275/8,402 | 4,401/8,403 | X coverage is 52.4% |
| ReBeat cognitive | 8,261/8,490 | 4,550/8,492 | X coverage is 53.6% |

IITB post-exercise participant-median all-method range changed by -3.82 ms for B and -31.70 ms for X. Because failures increased at the same time, the smaller retained-beat range cannot be called improved reliability. ReBeat X range increased by 15.59 ms (95% bootstrap interval 4.00 to 20.00; 17 of 24 participants increased).

# Physiological condition sensitivity

Matched detector analyses used the same 18 IITB participants at rest and post-exercise and the same 24 ReBeat participants at baseline and cognitive workload. IITB has no independent B/X reference. ReBeat expert-reference change was estimable for 22 participants represented in both conditions, with different beats recorded in each condition.

Figure 7. IITB paired B1 failure change. Each point is post-exercise minus rest for one participant; the diamond and interval show the median and participant-bootstrap 95% interval. B1 increased by 14.40 percentage points (95% interval 2.28 to 28.22). B7 showed the same median increase with a similar interval. This establishes condition sensitivity in detection success, not reference timing accuracy.

| Matched outcome | Median change | 95% interval or direction | Conclusion |
| --- | --- | --- | --- |
| IITB B1 failure | +14.40 points | 2.28 to 28.22 | Supported increase |
| IITB X4 failure | +7.73 points | -5.09 to 19.82 | Direction uncertain |
| IITB RR interval | -177 ms | Post-exercise minus rest | Strong condition change |
| ReBeat X8 failure | +0.63 points | 0.01 to 1.62 | Small supported increase |
| ReBeat X range | +15.59 ms | 4.00 to 20.00 | 17/24 increased |
| ReBeat B1 expert MAE | +2.00 ms | -2.00 to 4.00 | No consistent median change |
| ReBeat X12 expert MAE | 0.00 ms | 0.00 to 4.00 | No consistent median change |

Every one of the 23 ReBeat participant-median absolute expert-reference error intervals included zero, so the available annotations do not establish a consistent workload-related median accuracy change. Figure 8 aligns failure patterns across datasets without pooling. X4 is failure-prone in every dataset, but cohort, equipment, preprocessing and morphology are confounded, so the magnitude cannot be attributed to disease or a single physiological mechanism.

# Waveform morphology and failure investigation

All six prespecified features were evaluated against strict common-reference absolute error, attempted detection failure and method-specific peer disagreement for every applicable B and X definition. The 1,794 tests produced 23,631 participant effects. The table reports multiplicity-controlled exploratory support as positive/negative effects; the sign describes the measured association or failed-minus-valid contrast, not whether a waveform is physiologically better.

| Feature | Reference error | Failure | Peer disagreement | Observed pattern |
| --- | --- | --- | --- | --- |
| Sharpness | 12/69 (2+/10-) | 10/115 (0+/10-) | 21/115 (6+/15-) | Supported failure contrasts consistently had lower failed-beat sharpness; other outcomes remained method-specific |
| Local slope | 8/69 (3+/5-) | 10/115 (0+/10-) | 22/115 (6+/16-) | Supported failure contrasts consistently had lower failed-beat slope; reference-error directions were mixed |
| Amplitude | 7/69 (6+/1-) | 8/115 (2+/6-) | 15/115 (12+/3-) | Mostly positive reference-error and disagreement effects; failure was not a uniform low-amplitude pattern |
| Signal-to-noise ratio | 5/69 (3+/2-) | 7/115 (7+/0-) | 18/115 (12+/6-) | Supported failure effects were positive, contradicting a simple low-SNR explanation |
| Competing features | 10/69 (5+/5-) | 7/115 (0+/7-) | 26/115 (13+/13-) | Directions were mixed or counterintuitive; more counted features did not consistently mean worse reliability |
| Beat-to-beat variability | 0/69 | 6/115 (5+/1-) | 7/115 (6+/1-) | No supported reference-error association; selected failure and disagreement effects were mainly positive |

Support was uneven: HeartCycle contributed 13 B and one X association, IITB 42 B and seven X associations, and ReBeat 81 B and 55 X associations. This distribution, the mixed directions and the 1,595 unsupported rows rule out a universal morphology threshold. HeartCycle X12 competing-feature evidence did not survive correction, and beat-to-beat variability had no supported reference-error association for any method.

Outcome definitions matter. Error effects are within-participant correlations on valid strict common-reference beats. Failure effects compare the participant median feature on failed and valid attempted beats without assigning timing errors to failures. Peer-disagreement effects use common-valid beats and therefore remain sensitive to exclusion of difficult X beats.

## Selected participant aware effects and operational interpretation

Figure 6. Selected method-specific morphology effects. Left: HeartCycle B1 local slope versus absolute ECHO AVO error, one within-participant Spearman coefficient per participant (median rho 0.216, 95% bootstrap interval 0.117 to 0.305; q=0.0237; n=17). Right: IITB post-exercise X4 failed-minus-valid sharpness (median -460 dataset-specific curvature units, interval -1,002 to -381; q=0.00525; n=16). Diamonds and bars show the participant median and interval. These associations identify reliability boundaries but do not prove causation.

B1 searches for the last qualifying local minimum; B7 searches for the last second-derivative upward crossing. Their IITB post-exercise failures were dominated by no qualifying feature or crossing. B1 had 579 misses. B7 had 571 misses, comprising 564 no-crossing records and seven search windows with fewer than two samples, plus 19 order-invalid detections. Both participant-median failure rates increased by 14.40 percentage points. Failed beats had fewer counted competing features (median failed-minus-valid contrast -2; q=0.00396; n=17), while the change in competing-feature count correlated negatively with the change in failure (rho -0.82; q=0.0035; n=18). That direction is counterintuitive and does not establish a morphology mechanism.

X4 searches for the third-derivative minimum in a slice that includes C, then requires X to be strictly after C. Every recorded X4 failure was an order rejection because the minimum occurred at the excluded boundary. Lower failed-beat sharpness was reproducible in IITB post-exercise and ReBeat, but ReBeat failures also had higher amplitude and SNR and lower slope. The software mechanism is therefore established, while the physiological explanation remains multifactorial. X8 provides a contrast: its misses reflect absence of a qualifying post-minimum second-derivative maximum, but no failure-related morphology association survived correction; only selected disagreement associations were supported.

# Real waveform examples and paired agreement

Figure 1. HeartCycle failed or ambiguous X example selected after the full comparison. All markers belong to the same real beat. X4 was invalid because its selected sample did not satisfy strict C < X < next R; X8 missed because no qualifying post-minimum second-derivative maximum was found. The panel explains an execution path but is not population-level evidence. Clean, major-B and major-X cases for all datasets remain in the supplement.

Figure 9. HeartCycle X5 versus X6 agreement. Each point is a paired beat; the solid line is participant-balanced bias and dashed lines are limits of agreement. X5-X6 bias was +21.10 ms with limits -20.43 to +62.63 ms (17 participants; 1,706 beats). The related threshold methods are not interchangeable. Agreement between algorithms does not establish physiological accuracy.

# Integrated P1 findings and discussion

The evidence defines reliability profiles rather than a leaderboard. B8 had the lowest observed HeartCycle ECHO AVO median error and no failures there. B1/B7 had low ReBeat expert-label median errors but larger and condition-sensitive failure burdens. X12 combined a relatively low median error with a smaller HeartCycle tail than X2, while X4 exposed a boundary problem that removed 31–49% of attempted beats across datasets. These differences depend on reference provenance, tails, coverage, participant variation, condition and waveform context.

| Claim | Evidence | Boundary |
| --- | --- | --- |
| B8 had the lowest observed HeartCycle ECHO B median error | Figures 2–3; strict common sample | Descriptive comparison; HeartCycle ECHO only |
| X12 combined a low X median with a smaller tail than X2 | Figures 2–3; ECHO and expert tables | Reference types remain separate |
| X4 is operationally unreliable | Figure 4; exact failure records | Does not determine physiological correctness |
| IITB B1/B7 failure increases after exercise | Figure 7; matched n=18 | No independent IITB reference |
| ReBeat cognitive X disagreement increases | Figures 5 and 7; +15.59 ms | Success-only common-valid subset |
| ReBeat median expert-reference change is uncertain | Matched n=22; all 23 intervals include zero | Different beats; expert ICG reference |
| Morphology effects are method and outcome specific | Figure 6; 1,794-test screen | Exploratory and non-causal |
| X5 and X6 are not interchangeable | Figure 9; bias and limits | Agreement is not accuracy |
| Dataset differences cannot be disease effects | Figure 8; comparability audit | Acquisition and cohort confounding |

## Contribution beyond prior work

Árbol et al. compared three B algorithms against two scorers and documented large differences in missed cycles. Forouzanfar et al. developed a morphology-aware B detector and evaluated rest and post-exercise data. Karpiel et al. assessed B, C and X detections against expert labels in a clinical pilot. P1 does not repeat a single B ranking. It combines 10 B and 13 X definitions with independent ECHO evidence where available, expert-label evidence kept separate, exact failure-path reconciliation, matched condition changes, denominator-aware disagreement and an all-method six-feature morphology screen.

For downstream work, PEP inherits B timing error and LVET inherits both B and X error and their covariance. A low median error can coexist with severe tails, and a high-failure method can look stable on retained beats. P1 supports reporting the exact definition, reference type, coverage, signed bias, tails and condition-specific behaviour before calculating PEP or LVET.

# Limitations outstanding validation and conclusion

## Established findings

HeartCycle B8 and X12 had the lowest observed participant median absolute errors within their point-specific strict common-reference comparisons; paired superiority was not established.

Failure burdens and mechanisms differ sharply by method; X4 repeatedly selects an inadmissible boundary sample.

Exercise and cognitive workload change failure or disagreement for selected methods within matched participants.

Specific morphology features are associated with selected reference errors, failures and peer disagreement, but effects are heterogeneous.

## What remains uncertain

IITB has no independent B/X timing reference. ReBeat expert ICG labels are not interchangeable with ECHO valve events.

Common-valid analyses are success-only and particularly selective for X. Physical dZ/dt calibration is unverified.

No approved clinical timing tolerance exists. Morphology associations are exploratory and do not prove physiological causation.

The frozen detector, execution and filter source files are now packaged under src/variability_analysis and match the recorded SHA-256 hashes exactly.

Fresh raw-data replays reproduced the stored R, next-R, C and all canonical 10 B and 13 X detection outputs. An isolated Python 3.12 environment without system site packages reproduced all 84 IITB/ReBeat recordings and the full HeartCycle pipeline; raw datasets remain external.

## Direct answer to RQ1

HeartCycle B8 and X12 had the lowest observed median absolute errors against their respective ECHO events, while ReBeat expert annotations produced a different pattern. None is universally reliable. Method-specific failures, selective complete-case coverage, participant heterogeneity and condition changes determine when a timing estimate is dependable. Sharpness and slope often tracked failure or disagreement, but amplitude, SNR and competing-feature directions were inconsistent. The code-based failure mechanisms are reproducible; the waveform associations do not yet establish causal physiology.

## Reproducibility

The analytical package contains preprocessing, alignment, validation, statistics, morphology, figure-generation and report scripts, plus README, requirements and 23 passing tests. The evidence package includes the 115-row results table, all 1,794 morphology tests, 690 failure-reason rows, 672 condition-morphology tests, matched ReBeat expert-reference changes, plotting data, audit logs and the complete Figure 1 to 9 inventory. The three frozen upstream implementation files are now packaged and hash-verified. A full raw replay still requires the external raw datasets at the configured read-only paths.